# Yokogawa DL7400 Series Digital Oscilloscopes

# Sanitization Procedure

## Acquisition Memory

Acquisition memory, including channel and history memory, is 100% volatile memory and is completely erased every time the instrument is turned off. No additional sanitization is required.

## Internal Instrument States

The DL7400 has room for three internal instrument states (found in the SETUP ( STORE/RECALL menu). Follow this procedure to erase all internally saved settings:

- Turn OFF the Power switch.

- Turn the Power ON while holding down the RESET key. Hold down the RESET key until you hear a short beep.

- The instrument will continue to boot up. Once it is finished booting up, a message (code = 13) will appear on the screen. This is the indication that all instrument settings have been reset. You can confirm that all instrument states have been erased by going to the SETUP ( STORE/RECALL menu.

This procedure ensures that all settings, excluding the date / time setting, are initialized to factory default settings, including the setup data that have been stored using the store / recall function, the settings related to communications, the language setting, the USB keyboard language, and the SCSI ID setting.